PATENT OFFICE (PO), IIT JAMMU

INVENTION DISCLOSURE FORM (IDF)

# Title: Integrated Foldable Inflatable Neck Support Cushion with Regulated  
            Portable Oxygen Delivery System.

# Inventors:

| Sr. No. | Name | Nationality | Address |
| --- | --- | --- | --- |
| 1 | Anand Thakare | Indian | Department of Manufacturing Engineering & Industrial Management, COEP Technological University Pune. |
| 2 | Arjun P | Indian | Department of Mechanical Engineering, College of Engineering, Trivandrum. |
| 3 | Prajwal Jinde | Indian | Department of Aerospace Engineering, MIT ADT University Pune. |

# Abstract of the Invention:

The present invention relates to an integrated foldable inflatable neck support cushion with a regulated portable oxygen delivery system configured to provide simultaneous cervical support and supplemental oxygen to a user operating in high-altitude and oxygen-deficient environments. The invention comprises a foldable neck support cushion (1) having an outer cover fabricated from ripstop nylon (2) and an inner inflatable bladder constructed from 420D TPU-coated nylon (3). A detachable compressed oxygen canister (4) fabricated from Aluminium 6061-T6 alloy is securely mounted within a canister holding mechanism comprising internally threaded portions for threaded engagement with the canister and a lever-operated actuation mechanism configured to actuate the canister valve while the canister remains securely mounted. Upon actuation of the lever mechanism, compressed oxygen released from the canister is first passed through a primary pressure regulator that reduces the pressure from approximately 150 bar to about 5–8 bar. The regulated oxygen is directed into the inflatable bladder to inflate the neck cushion, thereby providing cervical stabilization and support. Oxygen subsequently exits the bladder through an outlet passage and passes through a secondary pressure regulator, which further reduces the pressure to a breathable level. The oxygen then flows through an adjustable flow regulator valve configured to provide oxygen flow rates of approximately 0–15 litres per minute before being delivered to the user through a flexible silicone oxygen delivery tube and a medical-grade silicone oxygen mask. The inflatable neck cushion functions both as a cervical support and as an intermediate oxygen flow chamber, enabling efficient utilization of oxygen for sequential cushion inflation and respiratory assistance. The integrated configuration eliminates the need for separate inflation devices and bulky oxygen carrying arrangements, thereby improving portability, rapid deployment, and ease of operation. The invention provides a compact, lightweight, foldable, and wearable emergency oxygen delivery system particularly suitable for military personnel, mountaineers, rescue teams, pilots, miners, emergency responders, industrial workers, and other users operating in oxygen-deficient or high-altitude environments. The invention enhances user safety by combining regulated supplemental oxygen delivery with cervical stabilization in a single integrated device.

# Background and Problem Statement:

# Personnel operating in high-altitude environments are frequently exposed to reduced atmospheric oxygen concentration. As altitude increases, the partial pressure of oxygen in the atmosphere decreases significantly, resulting in a condition known as hypoxia, which is characterized by an inadequate supply of oxygen to the body tissues. Hypoxia adversely affects physiological and cognitive functions, leading to symptoms such as breathlessness, dizziness, fatigue, reduced blood oxygen saturation, impaired physical performance, impaired decision-making, memory loss, hearing impairment, retinal damage, and, in severe cases, life-threatening complications.

# Military personnel deployed at high-altitude locations, mountaineers, rescue teams, pilots, miners, and other personnel operating in oxygen-deficient environments are particularly vulnerable to the effects of hypoxia. Although acclimatization procedures are followed before deployment, prolonged exposure to such environments may still result in reduced operational efficiency, impaired balance, and increased risk of injury. Medical management of hypoxia generally relies on portable oxygen systems or clinical oxygen therapy; however, such solutions are often bulky, require separate carrying arrangements, involve multiple setup procedures, and are not suitable for immediate emergency deployment under field conditions.

# Existing oxygen delivery systems are generally bulky, relatively heavy, and require dedicated carrying arrangements such as backpacks or separate carrying cases, making them inconvenient for continuous personal use and rapid emergency deployment. Such systems involve are not designed as compact, lightweight, wearable devices. Furthermore, these systems are intended solely for supplemental oxygen delivery and do not provide any form of cervical support or stabilization. Conversely, commercially available neck support cushions are designed exclusively to provide comfort or cervical support and do not incorporate any respiratory assistance or emergency oxygen delivery capability. 

Consequently, users operating in oxygen-deficient environments are required to carry separate oxygen delivery system which is not possible because of  increased equipment weight, reduced portability, longer deployment time, and an increased risk of neck instability during hypoxia-induced dizziness or disorientation.

# Furthermore, no known system integrates a foldable inflatable neck support cushion with a regulated oxygen delivery mechanism capable of simultaneously providing cervical stabilization and supplemental oxygen from compact, foldable and lightweight device. Existing devices fail to provide a compact, lightweight, easy to carry, foldable, and rapidly deployable solution that can be conveniently carried and operated in high-altitude or oxygen-deficient environments.

# Accordingly, there exists a need for an integrated, foldable, wearable device that combines emergency oxygen delivery with inflatable cervical support in a compact system, thereby enabling rapid deployment, improved portability, enhanced user safety, and simultaneous respiratory assistance and neck stabilization for personnel operating in oxygen-deficient environments.

# Detailed explanation of the invention along with working examples:

# The present invention relates to an integrated foldable inflatable neck support cushion with a regulated portable oxygen delivery system configured to provide simultaneous cervical support and supplemental oxygen to a user operating in high-altitude and oxygen-deficient environments. The invention combines a portable oxygen storage system, an inflatable neck support cushion, and a regulated oxygen delivery assembly into a single compact, foldable, and wearable device.

# The device comprises an outer cover fabricated from ripstop nylon enclosing an inner inflatable bladder constructed from 420D TPU-coated nylon. The inflatable bladder is configured in a generally U-shaped configuration adapted to support the cervical region of the user upon inflation.

# A detachable oxygen canister is provided as the oxygen source. In a preferred embodiment, the oxygen canister is fabricated from Aluminium 6061-T6 alloy and stores compressed medical-grade oxygen at a pressure upto 150 bar. The neck portion of the oxygen canister

# is provided with external threads for mechanical engagement with a canister holding mechanism.

# The canister holding mechanism comprises a housing having internally threaded portions at its lower end corresponding to the external threads of the oxygen canister, thereby 
enabling secure attachment of the canister by threaded engagement. The upper portion of the holding mechanism incorporates a lever-operated actuation mechanism positioned above the canister valve. The lever mechanism is configured to actuate the valve of the oxygen canister without requiring removal of the canister from the device. Upon actuation, the lever depresses or opens the valve mechanism of the canister, thereby initiating the controlled release of compressed oxygen.

# The released oxygen is first directed through a primary pressure regulator valve configured to reduce the pressure of the oxygen from approximately 150 bar to an intermediate pressure ranging from approximately 5 to 8 bar. The regulated oxygen at intermediate pressure is then conveyed through an inflation tube to the inner inflatable bladder. The inflatable bladder expands upon receiving the regulated oxygen, thereby forming a supportive cushion around the user's neck and providing cervical stabilization.

# An outlet tube is provided on the inflatable bladder for discharging oxygen after inflation. The outlet tube is connected to a secondary pressure regulator valve configured to further reduce the oxygen pressure from approximately 5–8 bar to a breathable pressure suitable for human respiration. Downstream of the secondary pressure regulator, a flow regulator valve is provided to adjust the oxygen flow rate according to the user's respiratory requirements. The flow regulator preferably provides an adjustable flow rate ranging from approximately 0 to 15 litres per minute.

# The oxygen leaving the flow regulator valve is conveyed through a flexible silicone oxygen delivery tube connected to a medical-grade silicone oxygen mask adapted to cover the nose and mouth of the user. The mask delivers supplemental oxygen at a controlled flow rate while the inflated cushion simultaneously supports the cervical region.

# The integrated configuration enables the oxygen stored in a canister to perform two functions, namely inflation of the neck support cushion and delivery of breathable oxygen to the user, thereby eliminating the need for separate inflation devices and large oxygen cylinder in high altitude regions.

# Working Example:

# In typical mode of operation, the device is maintained in a folded configuration during storage and transportation. During use, the user unfolds the neck support cushion and places it around the neck. The detachable oxygen canister is secured within the canister holding mechanism by engaging the external threads provided on the neck portion of the canister with the internally threaded lower portion of the holding mechanism. Once securely mounted, the user actuates the lever provided on the holding mechanism.

# Operation of the lever opens the valve of the oxygen canister, allowing compressed oxygen stored at pressure upto 150 bar to flow into the primary pressure regulator valve. The primary pressure regulator reduces the pressure to approximately 5–8 bar before directing the regulated oxygen into the inflatable bladder through the inflation tube.

# The inflatable bladder inflates to form a firm yet comfortable neck support cushion, thereby providing cervical stabilization to the user. Oxygen subsequently exits the bladder through the outlet tube and passes through the secondary pressure regulator, which further reduces the oxygen pressure to a breathable level suitable for inhalation.

# The breathable oxygen then passes through the adjustable flow regulator valve, allowing the user to select the required oxygen flow rate based on the prevailing environmental conditions and physiological requirements. The regulated oxygen is finally delivered through the silicone oxygen delivery tube to the medical-grade silicone oxygen mask, enabling the user to breathe supplemental oxygen while simultaneously receiving cervical support. After completion of use, the user closes the oxygen supply, allowing the cushion to deflate. The oxygen canister may then be unscrewed from the holding mechanism and replaced with a fresh canister when required. The device can subsequently be folded into a compact configuration for storage or transportation.

# Aspects that you want to monopolize:

## An inflatable neck support cushion configured as an intermediate oxygen flow chamber, wherein regulated oxygen first inflates the cushion to provide cervical support and is subsequently discharged through an outlet passage for delivery to an oxygen mask.

## A wearable, foldable inflatable neck support cushion integrated with a portable oxygen delivery system in a single compact device.

## A canister holding mechanism comprising a housing having internally threaded portions for threaded engagement with a detachable oxygen canister, the housing further incorporating a lever-operated actuation mechanism configured to actuate the canister valve while the canister remains securely mounted within the housing.

## A regulated oxygen delivery pathway extending from the oxygen canister through a primary pressure regulator, the inflatable neck cushion, a secondary pressure regulator, a flow regulator valve, an oxygen delivery tube, and an oxygen mask to provide controlled supplemental oxygen to the user.

## A compact, lightweight, foldable, and wearable configuration enabling rapid deployment and simultaneous cervical stabilization and supplemental oxygen delivery for use in high-altitude and oxygen-deficient environments.

# Details of the disclosure if invention has been disclosed to anyone outside Institute:

## The invention has not been disclosed to any person or organization outside the   
       institute and remains confidential.

# What is the Technology Readiness Level (TRL) of your invention?

# Technology Readiness Level 4 (TRL 4).

# Has your invention being reduced to practice?

# No. The invention has not yet been reduced to actual practice. The engineering design, working principle, and prototype development have been completed.

# Brief description on Utility & Market Potential:

Utility:

The invention provides a compact, foldable, and wearable emergency oxygen delivery system integrated with an inflatable neck support cushion for simultaneous supplemental oxygen delivery and cervical stabilization. It enables rapid deployment in oxygen-deficient environments and is intended for use by military personnel, mountaineers, rescue teams, pilots, miners, emergency responders, and industrial workers.

# Market Potential:

# The invention has significant commercial potential in the defence, healthcare, aerospace, mountaineering, mining, industrial safety, and disaster management sectors. The increasing demand for portable emergency oxygen systems and wearable safety equipment, particularly for high-altitude and oxygen-deficient environments, offers substantial opportunities for commercialization in both domestic and international markets.

# Market Players:

Oxy99
Linde Healthcare  
Drager  
Mountain High Oxygen  
Meggitt Aerospace Oxygen Systems  
Aerox Aviation Oxygen Systems

# Brief description on Novelty:

# The novelty of the present invention lies in the unique integration of a foldable inflatable neck support cushion with a regulated portable oxygen delivery system in a compact, lightweight, and wearable device. The invention incorporates a canister holding mechanism comprising a housing having internally threaded portions at its lower end for threaded engagement with the external threads of a detachable oxygen canister, thereby securely retaining the canister within the device. The upper portion of the housing incorporates a lever-operated actuation mechanism configured solely to actuate the canister valve and initiate the controlled release of oxygen from the canister while the canister remains securely mounted in the holding mechanism. The inflatable neck cushion performs a dual function by providing cervical support and acting as an intermediate oxygen flow chamber before oxygen is delivered to the user.

# A dual-stage pressure regulation arrangement ensures safe inflation of the neck cushion and controlled supplemental oxygen delivery through an oxygen mask. This integrated configuration enables simultaneous cervical stabilization and respiratory assistance with rapid deployment, making the invention particularly suitable for high-altitude and oxygen-deficient environments.

# Brief description on Non-obviousness/Inventive step:

The present invention involves an inventive step by providing a technically integrated wearable system that performs simultaneous cervical support and regulated supplemental oxygen delivery using a single portable oxygen source. Unlike conventional oxygen delivery systems and neck support devices, which are designed and operated independently, the present invention employs the oxygen stored within a detachable compressed oxygen canister to sequentially inflate an inflatable neck support cushion and thereafter deliver breathable oxygen to the user through a controlled oxygen delivery pathway.

A significant inventive feature of the invention is the canister holding mechanism comprising a housing having internally threaded portions for threaded engagement with the detachable oxygen canister, while the upper portion of the housing incorporates a lever-operated actuation mechanism configured solely to actuate the canister valve without removing the canister from the housing. This arrangement enables secure retention of the canister together with rapid activation of oxygen supply, thereby improving operational reliability and ease of deployment under emergency conditions.

The invention further incorporates a dual-stage pressure regulation arrangement in which a primary pressure regulator reduces the high-pressure oxygen from the canister to an intermediate pressure suitable for inflation of the neck support cushion, while a secondary pressure regulator further reduces the pressure to a safe breathable level before controlled delivery through an adjustable flow regulator valve and oxygen mask. This sequential utilization of the same oxygen source for both cushion inflation and respiratory assistance provides a technical solution that is not achieved by conventional portable oxygen systems.

The coordinated integration of the foldable inflatable neck support cushion, detachable oxygen canister, threaded canister holding mechanism, lever-operated valve actuation mechanism, dual-stage pressure regulation system, and regulated oxygen delivery pathway produces a compact, lightweight, wearable, and rapidly deployable emergency support device. The present invention provides a technical advancement over the existing technology through the specific functional integration of its structural and operational features. The coordinated interaction of these features is not rendered obvious by the known prior art and provides technical advantages that would not have been readily apparent to a person skilled in the art at the time of the invention.

# Prior Art/References:

### Vento Portable Hypoxia Device

The Vento device is a portable solution designed for emergency treatment of hypoxia. It utilizes manual air pumping to deliver positive pressure on the chest, which helps to improve oxygen saturation by simulating lower altitude atmospheric conditions. However, the Vento device does not supply supplemental oxygen directly and instead relies on ambient air manipulation. Furthermore, it does not incorporate any neck support functionality, requires continuous manual operation, and does not provide hands free emergency use capability.

### Summit Oxygen Systems

Summit Oxygen Systems provide supplemental oxygen using high pressure, refillable cylinders with specialized regulators and dedicated masks. While effective for oxygen delivery, these systems are bulky, complex, and not designed as wearable & portable devices. They require significant manual setup, lack ergonomic integration with the body, do not provide any cervical support, and are not optimized for rapid emergency deployment in field conditions.

### Standard Portable Oxygen Kits

Conventional portable oxygen kits consist of separate oxygen cylinders, regulators, tubing, and masks assembled as discrete components. These kits require multi step assembly, do not offer wearable neck support, lack foldable or compact form factors suitable for personal carry, and do not integrate oxygen delivery with cervical stabilization.

### Travel Neck Pillows

Commercially available travel neck pillows provide cervical support through foam, inflatable bladders, or memory foam constructions. However, these are purely comfort devices with no respiratory function. They do not incorporate oxygen delivery capability, lack regulated gas management systems, and are not designed for emergency medical applications.

### Military Emergency Oxygen Systems

Military grade emergency oxygen systems are typically designed as standalone backpack units or fixed installation systems. These are heavy, require dedicated carrying arrangements, and do not integrate neck support functionality. Their size and weight make them impractical for individual carry during tactical operations.

### Medical Oxygen Backpacks

Medical oxygen backpacks combine oxygen cylinders with carrying frames and delivery tubing for mobile medical applications. While portable, these systems are designed for medical personnel rather than individual users, do not provide cervical support, require dedicated backpack carrying arrangements, and are not optimized for emergency self administration.

### Portable Oxygen Concentrators

Portable oxygen concentrators extract and concentrate oxygen from ambient air through molecular sieve technology. These battery powered devices provide continuous or pulse dose oxygen delivery but require electrical power, are relatively heavy, do not provide neck support, and are not designed for rapid emergency deployment in harsh field conditions.

# Deadlines/ Due dates:

- Is there an approaching deadline for protecting the invention from public disclosure?

NO.

- Please specify the deadline:
Not Applicable.

# Potential Licenses:

Table 1.

| Organization | Contact Person | Email | Contact Number |
| --- | --- | --- | --- |

None to declare.